AEI CALIBRATION DIRECTIONS

- Build mouth piece

-white is inspired side – connects to turbine

-clear is expired side – connects to sensors

- Turn on computer

- Open “NewVO2” program by double clicking the icon on the desktop

- Turn on turbine – KL engineering (tall cart system by front column) – Vacumed (small cart system by back wall)

- Turn on pump

Calibration

Follow prompts, after completion of each section “done” icon should be in green.

- Subject Data

-enter name, age, height, weight

-click return

2.   Environmental & Peripheral conditions

-enter barometric pressure – check weather station.  Should be around 620-640mmHg

-enter temperature and humidity – use either WBGT (Celsius) or sling cyclometer (Fahrenheit)

-click return

Note: this program sometimes recalls itself a second time.  Just clear the initial prompts and the data will still be there and click return again!

- Calibrate single turbine

-click the turbine system you’re calibrating either KL or Vacumed

-get out 3L syringe

-connect turbine to syringe (use rubber connector for KL engineering or directly to syringe for Vacumed).  Remember, small whit ends receive air flow!

-click collect signal

-perform 6 full syringe pumps – be gentle – results should be around 0.3v for KL and around 6v for Vacumed

-click collect signal

-delete row zero, and any other row that has an incorrect total voltage

-click accept value

Note: this program sometimes recalls itself a second time.  Unlike the Environmentals sub-program, the prior data is removed and you have to re-calibrate the turbine!!!!!

Place the turbine back in its protective box after calibration!

- Calibrate gas analyzers

-you will be measuring calibration gas, nitrogen, and room air

-select Room Air

-make sure sample line is in open to room air.  I recommend that you still place the sample line in a calibration syringe to make sure no-one breaths over it!

wait until graph lines stabilize to the sampled gas values – this should be easy for the room air calibration condition, as this gas condition is pre-existing in the syringe when you do this first.

-click “click to collect” signal bar –

-click “collecting: click to stop” (signal bar again) to end collection after about 5-10 s of collection.

-select Cal Gas

-place sample line inside calibration tank syringe

-turn on pressure gauge – turn on gas flow – listen for gas

- wait until graph lines stabilize to the sampled gas values

- click “click to collect” signal bar

- click “collecting: click to stop” (signal bar again) to end collection after about 5-10 s of collection.

-select Nitrogen

-place sample line inside nitrogen tank syringe

- turn on pressure gauge – turn on gas flow – listen for gas

- wait until graph lines stabilize to the sampled gas values

- click “click to collect” signal bar

-graph should zero out for both analyzers

- click “collecting: click to stop” (signal bar again) to end collection after about 5-10 s of collection.

-once done calibrating click “calculate regression”

-you should see linear graphs with data points

-oxygen should have 3 data points

-carbon dioxide should have 2 points close to zero and one at ~5%

If the O2 calibration linear regression does not fit all three data points, then you need to repeat the calibration.  To do this, click “accept value”, and re-enter the sub-program again from the main menu list.

-click accept value when finally happy with your calibration results

- Calculate delay time

-using either Nitrogen or Cal Gas to measure delay time in sample line

-place sample line in syringe

-once analyzer reading is stable click stable value

-simultaneously you want to turn gas on and click to start

-values should be around 3.5-4.5s for KL system and 2-3s for Vacumed

-click stop to exit sub-program, or “reset” to run the routine again.

- Done with gas

-bleed gas lines and flow regulators after turning tanks off

- Check signals

-connect mouthpiece to turbine (inspired side) and mixing bag (expired side)

-connect sample line to mixing  bag

-have subject breath into it to check signal - remember nose clip

-click continue

- Collect data

-run test – have subject already warmed up based on prior lab protocol development procedures

-click start test with the subject at rest.  Look at data at rest to check that all is OK.  Look at RER, VO2 and VE.

-RER between 0.7-1.0

-VO2 ~3.5 – 5ml/kg/min STPD

-VE ~6-8 L/min STPD

-start protocol after about 1-2mins of rest collection

-click stop test when done – record time when test was terminated

- Turn off turbine and pump

Use “Process data” sub-program to process data and save a txt file of processed data.  Make sure you know the path of where this file gets saved!!!

- The initial saved data file is automatically saved to the “data” folder of the NewVO2 program.  There is a link to this folder on the desktop.

- Clean mouth piece.  To do this properly, do as follows.

- Wash all components to be clear of saliva under running tap (faucet) water

- Dis-assemble the mouthpiece

- Place all components, except for the fragile one-way valves, into the Cydex sterilization solution.  Use gloves to avoid contact the cydex solution.

- Place the washed one-way valves on the drying rack.

- Components need to be sterilized for 10 min.

- After 10 min, rinse in tap water and place on drying rack.

- Rinse expired mixing bag thoroughly with water and let dry on rack

- Check all AEI system components before you leave.  Turbine should be off, and flow pump should be off. Gas analyzers always remain on!